7600 SERIES SCALE - OPERATIONS - Section 1

CONTENTS

INTRODUCTION - 7600 Series Scale			3

INSTALLATION and  SET-UP				4

DIAGNOSTICS						5

OPERATION

OPERATIONAL FEATURES			6

EXTERNAL CONNECTORS				6

REMOTE BASE PIN-OUT			7

DISPLAYS AND INDICATORS			7

PUSH BUTTON FUNCTIONS			8

WEIGHING

WEIGHING WITHOUT TARE			10

WEIGHING WITH TARE				10

COUNTING

NORMAL COUNTING (one switch)			12

WITH SAMPLE 				12

WITH KEYPAD PIECE WEIGHT

ENTRY				13

ALTERNATE COUNTING METHODS		14

TWO SWITCH COUNTING			14

TOP END COUNTING (one switch)		14

AUTOMATIC SAMPLE-to-BULK		14

DETERMINING ERROR OF COUNT

(one switch)				14

DETERMINING ERROR OF COUNT

(two switch) 				15

PIECE WEIGHT DISPLAYED AS

WEIGHT/1000 pcs.			15

NEGATIVE COUNTING			15

NORMAL COUNTING (two switch)		15

NORMAL COUNTING WITH

AUTOMATIC SAMPLE-to-BULK	16

CONTENTS (cont’d.)

COUNTING (cont’d.)

ALTERNATE COUNTING METHODS

TOP END COUNTING

WITH SAMPLE

(one switch)				17

TOP END COUNTING

WITHOUT SAMPLE

(two switch)				17

ERROR COUNT USING AUTO

SAMPLE UPDATE			18

ERROR COUNT (two switch)		18

ERROR COUNT WITH AUTO

SAMPLE-BULK

(two switch)				19

NEGATIVE COUNTING

WITHOUT SAMPLE

(two switch)				19

BI-DIRECTIONAL SERIAL ASCII

INTRODUCTION					20

COMMUNICATION CAPABILITIES		20

RS-232 CONNECTOR PIN-OUT			21

RS-232 COMMUNICATIONS PARAMETERS	21

COMMAND FORMATS				22

SCALE COMMANDS				23

SAMPLE COMPUTER PROGRAM			25

ERROR MESSAGE LIST					27

7500 SERIES SCALE - PROGRAMMING/CALIBRATION -Section 2

CONTENTS

INTRODUCTION							1

CALIBRATION							2 - 9

FORMATTED PRINT						10 - 11

FORMATTED PRINT CODES					12 - 13

COMMAND FORMATS						14

RS-233 SCALE COMMANDS					15 - 16

RS-232 COMMUNICATION PARAMETERS			17

RS-232 PRINT CODES						18

SLEEP MODE and TIME/DATE					19

RS-232 SERIAL INTERFACE					19 - 20

CALIBRATION COMMANDS					21 - 22

CALIBRATION EXAMPLE						23

“CAL” COMMAND							24 - 26

7500 SERIES SCALE - OPERATIONS

SECTION 1

INTRODUCTION

This manual covers the Operations and Programming/Calibration of the Pennsylvania Model 7600 Series Scale.

The 7600 Scale has been designed to be simple and easy to operate.  It  also has many features that make it extremely versatile to use such as:

-  Standard RS-232 Interface with a selectable output for

communications with a computer or printer.

-  Adjustable response time

-  Single board electronics for ease of servicing.

-  Simple one button counting operation.

-  Weight reading may be switched between two weighing units.

-  May be programmed as either weigh and count or weigh only.

-  Keypad entry of tare weight.

-  Keypad entry of piece weight.

-  Optional second weighing platform.

In addition, the 7600 Scale can be calibrated with the keypad or through the RS-232 Interface to match the requirements of the application on hand, as it employs a digital calibration technique that eliminates dip switches and potentiometers.

INSTALLATION and SET-UP INSTRUCTIONS

The 7600 Series Scale has been packaged for shipment to ensure safe, damage-free arrival.  Please use reasonable care when removing the unit from the shipping carton, and retain the original packaging materials in case reshipment is required.

NOTE:  	TO PROTECT THE WARRANTY, FILL OUT AND RETURN

THE WARRANTY CARD AND THE SCALE CONDITION REPORT.

To prepare the 7600 for operation, follow the procedure outlined below.

1. 	After opening the shipping carton, remove the molded foam top from the carton. (On 2 lb. and 5 lb. capacity scales the platform is packaged on top of this foam, remove it first and lay it aside.)

2. 	Gently lift and remove the stainless steel platform cover only. (On 2 lb. and 5 lb. capacity scales the platform is packaged separately and does not need to be removed.)

3. 	Remove any options which may be packed with the scale.

4. 	Carefully remove scale from the packaging by grasping both sides of the base.

DO NOT LIFT SCALE BY THE TOP SPIDER OR SUB-PLATFORM ASSEMBLY.

5. 	Place the scale on a stable, level surface for operation.

6. 	Adjust the corner leveling feet until the level bubble indicates the unit is level. Firmly tighten hex jam nuts on the leveling feet. (Any time the scale is relocated, it should be leveled.)

7. 	Remove the protective plastic wrap from the platform and place the platform on the spider.

8. 	To activate the scale, plug the line cord into any grounded 50/60 hertz 120 volt outlet. The scale will first display the version number of the software installed followed by its diagnostic countdown process.  During this countdown the display will show 9.9.9.9.9.9. through 0.0.0.0.0.0.  each number representing a self-diagnostic test to verify the operation of important unit functions. If the scale should fail any of these tests the display will freeze on that number and display an error message, refer to page 23 for a list of error messages. (The diagnostic failure may be bypassed by pressing the CHECK button.) The serial interface on all models will also transmit the test numbers. The following table lists each test and its corresponding display.

DISPLAY				DIAGNOSTIC FUNCTION

9.9.9.9.9.9.				EPROM Checksum, light all LEDS

8.8.8.8.8.8.				RAM check

7.7.7.7.7.7.				Novram check

6.6.6.6.6.6.				Check switches not shorted

5.5.5.5.5.5.				No Test

4.4.4.4.4.4.				Check load cell excitation

voltage

3.3.3.3.3.3.				Analog Verify (checks output

from a-d)

2.2.2.2.2.2.				Load CFG EEPROM, Test Virgin,

Checksum, Initialize on error

1.1.1.1.1.1.				Load CAL EEPROM, Test Virgin,

Checksum, Initialize on error

0.0.0.0.0.0.				Test RS-232 connections

-  -  -  -  -  - 				Initializing system (may take

several seconds)

NOTE:	Allow at least 20 minutes for initial warm-up and load cell stabilization.

- The system should now be ready for operation.

OPERATION

OPERATIONAL FEATURES

The following features are available on the 7600 Series Scales.

Overload  -  If the weight exceeds 105% of full capacity, the scale will read out                                       "OLOLOL", meaning overload.

Underload  -  If the weight is below -3% of full capacity, the scale will read out                                         "ULULUL", meaning system underload.

Automatic Zero-Tracking (AZT) - Maintains the system zero and is normally set to +/-                                      1/4 of the display resolution ( The resolution being                                        the smallest increment the scale is programmed to                                         read). AZT increases the overall accuracy of the                                           system by removing errors at zero caused by                                                 accumulation of debris on the base/container or                                             minor drifts in the electronics or load cell.

Count/Weigh  -   Scale is capable of both counting and weighing functions. The count                                    may be turned off in calibration if not needed.

Four Selectable Sample Sizes  -   Four different sample sizes may be entered into the                                                               scale memory during programming.

Counting Functions  -  A variety of different counting methods may be employed,                                                   depending on type of counting selected during calibration.

Sample Update  -   Piece weight accuracy may be improved after the initial sample by                                        adding a quantity of parts between 10% and 100% of the original                                        sample to the platform. When this is done the scale will recalculate                                      the piece weight, display the new percent error/accuracy for 2                                              seconds and then return to the counting mode.

Second Base  -  An optional second base is available to expand the capacity of the scale.

EXTERNAL CONNECTORS

The following connectors may be accessed through the opening on the bottom of the scale.

D-Subminiature Connector (9-pin)  -   Used with RS-232 Interface.

EXTERNAL CONNECTORS (cont’d.)

Remote Display Connector  -  Used to connect a remote display option to the scale.

REMOTE BASE PIN-OUT

Round CPC Style connector (8 pin)  -  Only found on scales with second base option                                                           installed. This connector is the mating                                                                       connector for Pennsylvania remote bases, if the                                                        7600 is to be attached to a base of another                                                                manufacture a mating connector kit, part number                                                      48445 may be ordered from your Pennsylvania                                                         Scale distributor. For wiring purposes the                                                                 following chart provides the pin out for the                                                               Remote Base Connector:

PIN NUMBER			FUNCTION

1 ------------------------------------------ +	Signal

2 ------------------------------------------ -	Signal

3 ------------------------------------------  	Ground

4 ------------------------------------------  	Not Used

5 ------------------------------------------ -	Sense

6 ------------------------------------------ -	Excitation

7 ------------------------------------------ +	Sense

8 ------------------------------------------ +	Excitation

DISPLAYS and INDICATORS

Weight Display  -  Indicates weight when the scale is weighing. Indicates count when                                       scale is in the counting mode, and displays messages during                                                 calibration and data entry operations.

Gross Indicator  -  Illuminates when the weight displayed is the Gross weight.

Tare Indicator  -  Illuminates when ever a tare weight is entered in memory.

Net Indicator  -  Illuminates when the weight displayed is the Net weight.

Zero Indicator  -  Illuminates when the weight is within the programmed limit of system                       zero. (Normally 1/4 of one display resolution.)

DISPLAYS and INDICATORS (cont’d.)

Stable Indicator  -  Illuminates when the weight on the display is stable within the                                              programmed limit. (Normally 1/4 of one display resolution.)

Count Indicator  -   Illuminates when the scale is in the count mode.

Secondary Weighing Unit Indicator  -  Illuminates when the weight displayed is in the                                                                     secondary weighing unit.

Primary Weighing Unit Indicator  -  Illuminates when the weight displayed is in the                                                                      primary weighing unit.

Base 2 Indicator  -   Illuminates when the optional second base has been selected.

PUSH-BUTTON FUNCTIONS

ZERO Button  -   Sets the scale to zero.

SAMPLE SET Button  -   Places the scale in the sample set mode for counting                                                             operations.

GROSS/NET Button  -  Switches the scale display between Gross weight and Net                                                     weight.

PIECE WEIGHT Button  -  Will cause the scale to display the presently stored piece                                                       weight, if any, and puts it into the piece weight entry mode.

TARE Button  -  Will tare off any weight setting on the platform and enter it into the tare                               memory.

KEYPAD TARE Button  -  Will display whatever tare is presently stored in the tare                                                        memory and puts the scale into the keypad tare entry mode.

BASE/1 Button  -  Used to switch the scale between base 1 and base 2, when the second                                   base option is installed.

%/3 Button  -  Used to display the percent of error, or percent of accuracy. (Only if these                           features have been selected during programming.)

UNITS/4 Button  -  Used to switch the scale between the primary weighing unit, secondary weighing unit, and the count mode.

PUSH-BUTTON FUNCTIONS (cont’d.)

ID/5 Button  -  Will display whatever ID number is presently stored in the ID memory                                  and puts the scale into the ID entry mode.

CHECK/6 Button  -  Pressing this button will cause the scale to initiate the diagnostic                                          countdown. It may also be used as a continue key if the scale                                               should freeze on any of the diagnostic numbers during countdown.

PRINT/7 Button  -  Transmits formatted print to any device connected to the RS-232                                         port.  Holding this push-button for four seconds will access the                                            RS-232 configuration mode.

./CLEAR Button  -  Used when the scale is in one of the data entry modes. Pressing this                                     button once will put a decimal point onto the display and pressing it                                     a second time will clear the display.

ENT Button  -  Used when the scale is in one of the data entry modes. Button is used to                              enter information into the system memory.

WEIGHING

WEIGHING WITHOUT TARE

If the second base option has been installed, press the BASE button to select base 1 or base 2. When base 2 is selected the base 2 LED will be lit. When base one is selected the base 2 LED will be out. If base 2 has been locked out during calibration, the word "Error" will momentarily show on the display.

Press the UNITS button to select either the Primary weighing unit or the secondary weighing unit. (Associated indicator will be lit.)

Verify that the "GROSS"  indicator is lit. If not, press the GROSS/NET button to light this indicator and put the scale into the gross weight mode.

Establish a base zero by pressing ZERO with nothing on the scale platform to clear any existing weight readings.

Place the item(s) to be weighed on the platform and read the weight on the display.

NOTE:  	If very light items (less then 1/4 of the display resolution) are placed on the platform individually, the weight may be zeroed off by the AZT feature.  Add light items to the platform simultaneously.

WEIGHING WITH TARE

If the second base option has been installed, press the BASE button to select base 1 or base 2. When base 2 is selected the "base 2" LED will be lit. When base one is selected the base 2 LED will be out. If base two has been locked out during calibration, the word "Error" will momentarily show on the display.

Press the UNITS button to select either the Primary or the Secondary weighing unit. (Associated indicator will be lit.)

Press the ZERO button with nothing on the platform to establish a base zero.

WEIGHING WITH TARE (cont’d.)

A tare weight may be acquired using one of the following methods:

1.)	Place the container or object to be tared off on the

platform and press TARE; or

2.) 	Press the KEYPAD TARE button and then key in the weight

of the container or object, and press ENT.

The net weight value will be displayed on the weight display and the NET indicator will be lit.

Pressing  the GROSS/NET button will toggle between gross weight and net weight.

When a tare value is stored in the system, the TARE indicator will be illuminated.

Place the objects to be weighed in the container and read the net weight on the display.  To clear a tare value, remove all weight from the scale and press TARE, or press KEYPAD TARE, key in 0 and press ENT.

NOTE: 	IF the tare has been acquired using the TARE button, the tare value may be viewed by pressing KEYPAD TARE. Pres ENT to return to the weight display without altering the tare weight.

COUNTING

An electronic counting scale counts parts by knowing the weight of a single piece, dividing that into the total weight on the scale and displaying the result. With the 7600, this piece weight may be entered into the scale by one of two methods:

FIRST:	Counting out a "sample" quantity of the parts and placing them onto the scale. The scale will take the weight of the sample, divide it by the known quantity and derive a piece weight used to count the rest (or "bulk") of the parts.

SECOND: 	If the individual piece weight of the parts is already known, this weight value may be entered into the scale via the keypad.

SAMPLE SIZE:

If the "sample" method of piece weight entry is used, the total weight of the sample must be at least 0.04% of platform capacity or the scale will not recognize the sample. When sampling, it is always advantageous to use the largest sample possible. Not only does the larger quantity tend to average variations in the weight of individual pieces, but the heavier sample weight improves the counting accuracy of the scale itself.

NORMAL COUNTING PROCEDURE

ONE SWITCH COUNTING  -  The 7600 scale is normally shipped from the                                                             factory set up to follow this counting procedure.                                                      Unless your scale has been programmed for one of                                                   the alternate counting choices,  this procedure is                                                       the only one necessary to learn.

NORMAL COUNTING - Using Sample

- If a container will be used to hold items being counted, place it on the platform.

2.	Press the SAMPLE SET button. The scale zeroes. "AddXXX" appears on the display (where "XXX" is one of the four pre-programmed sample sizes.). Repeated pressing of the switch successively displays the four sample sizes entered during calibration.

- If a different sample size is required, key in the desired sample size (do not press ENT).

4.	Place the entire sample on the platform at one time. (If the scale has been programmed to show the percent of error; this value will be displayed momentarily at this time.) The scale will now  automatically switch to the count mode, displaying the number of parts in the sample.

NORMAL COUNTING  - Using Sample (cont’d.)

- All of the remaining parts may now be added to the scale and counted. ZERO, TARE and KEYPAD TARE buttons may now be used without affecting the piece weight.

6.	To count a different part, return to step one.

NORMAL COUNTING - Keypad Piece Weight Entry

If the individual weight of the parts being counted is known in advance the piece weight may be entered through the keypad. This eliminates the need for the sample set operation. (If the percent of error feature has been selected in programming, the scale will not calculate the percent of error when keypad entry of piece weight is used.)

- Select the appropriate weighing unit (primary or secondary) for the piece weight being entered by pressing the UNITS/4 switch.

- Press the PIECE WEIGHT switch. The display will alternately show "PC" and any previously stored piece weight.

3.	Key in the weight of one piece and press ENT.

- The scale will go into the count mode.

5.	If a container will be used to hold the parts, place it onto the scale and use the ZERO and/or TARE buttons to zero the scale. Add the parts to the scale at this time. NOTE: The parts in a full container may also be counted by placing the full container onto the scale and using KEYPAD TARE to key in the tare weight of the container.

ALTERNATE COUNTING METHODS

The 7600 may be programmed to perform a wide variety of different counting methods. the following section describes these counting methods, refer to the "Calibration and Programming Manual" for information on setting your scale to perform one of these.

Two Switch Counting

The 7600 can be programmed so that when the scale is in the sample set mode, the sample will not be accepted until the ENT button is pressed. This allows parts to be added to the platform one at a time, where as with the one switch method all of the sample must be added at once.

Top-End Counting

The total number of items in an unsealed container can be determined without removing the items from the container (except a sample). The container with parts is placed on the platform and the piece weight is calculated when a sample quantity is either added to (one switch) or removed from (two-switch operation) the container. Based on the piece weight, the total number of items in the container is displayed.

Automatic Sample-to-Bulk Counting

This method requires that the scale be configured with two bases, a "sample" and a "bulk" base. The lighter capacity "sample" base is used to calculate the piece weight, and the parts are counted on a heavier capacity "bulk" base. In most cases, higher counting accuracy is achieved with this method because the calculated piece is more precise when the sample is weighed on the light-capacity sample base. In normal operation, when the sample is placed on the sample base, the scale automatically switches to the bulk base and items to be counted are placed on the bulk base.

Counting by First Determining Error of Count (One-Switch Method using Automatic Sample Update)

The percent error/accuracy of count can be updated by adding more items to the sample after the initial sample size is placed on the platform and the percent error/accuracy has been displayed. In normal operation , the sample is placed on the platform and a percent error/accuracy is momentarily displayed, followed by the count. More pieces are then placed on the platform in the range of 10% to 100% of the original sample. A new piece weight is calculated based on the larger sample size, and the new percent of error/accuracy will be momentarily displayed. This may continue until a satisfactory percent of error/accuracy is achieved.

ALTERNATE COUNTING METHODS (cont’.d)

Counting by First Determining Error of Count (Two-Switch Method)

The percent error of count can be determined before the piece weight is established. Sample pieces are added to the platform until the percent of error/accuracy is acceptable. The sample size (number of pieces) is then keyed-in and entered, and the scale is ready to be used for counting.

Piece Weight Displayed as Weight Per 1000 Pieces

If selected in calibration, the piece weight can be displayed as weight per 1000 pieces. This option is useful when the items being counted are very light and the piece weight for one item would be an extremely small number. When entering the piece weight through the keypad (for normal counting), enter the weight for 1000 items if the piece weight is to be displayed as weight per 1000 pieces.

Negative Counting

This method is used to count the number of items removed from a container without first knowing the piece weight. The number of items removed from the container is shown on the display as a negative number.

Normal Operation with Two Switch Counting

- If a container will be used to hold items being counted, place it on the platform.

2.	Press the SAMPLE SET button. The scale zeroes. "AddXXX" appears on the display (Where "XXX" is one of the four pre-programmed sample sizes.). Repeated pressing of the switch successively displays the four sample sizes entered during calibration.

- If desired, a different sample size may be keyed into the scale. As this number is entered, the word "Add" will disappear and only the count indicator will be lit to indicate that the scale is in the sample entry mode.

4.	The sample may now be placed onto the scale, all at once or one at a time. (If the percent of error feature has been selected in programming, the percent of error will be displayed as the parts are placed onto the platform. The more parts that are added, the lower the percent of error will be.)

- Press the ENT button. The scale will switch to the count mode and display the sample count.

6. The remaining parts to be counted may now be placed onto the scale.

Normal Operation with Automatic Sample to Bulk Feature

The second base option needs to be installed and calibrated for automatic sample to bulk to operate. Base one is always the sample base and base two is always the bulk base.

1.	Press the BASE/1 switch to select the bulk base. (Base two indicator should be lit)

- Press the ZERO switch to zero the bulk base.

3.	If an empty container will be used to hold items on the bulk base place it on the platform and press either the ZERO or TARE switch.

- If a full container will be counted on the bulk base. Place the container on the platform, press the KEYPAD TARE switch and key in the weight of the container, press ENT.

5.	If a container will be used to hold the sample on base 1, place the container on the scale at this time.

- Press the SAMPLE SET button. The scale will switch to the sample base (base 1). "AddXXX" will appear on the display (where "XXX" is one of the four pre-programmed sample sizes). Select the desired sample size by repeatedly pressing the SAMPLE SET switch, or key in an alternative sample size.

7.	If the scale is programmed for a one-switch operation place the entire sample on the sample base all at one time. The scale will momentarily display the percent error/accuracy (If selected in calibration.), switch to the bulk base and enter the counting mode. The bulk base is now ready to count the parts.

- If the scale is programmed for a two switch operation, place the sample parts onto the platform all at once or one at a time. The display will show the percent of error/accuracy (If selected in calibration.). When all of the sample parts are on the scale press ENT. The scale will switch to the bulk base and enter the count mode. The bulk base is now ready to count the parts.

9.	At this point ZERO and TARE switches may be used without affecting the piece weight.

Top-End Counting (One-Switch Method; Sample Added)

1.	Press ZERO to zero the scale.

2.	Place container with parts on the platform.

3.	Press KEYPAD TARE, key in the tare weight of the container, then press ENT.

4.	Remove the sample from the container.

- Press the SAMPLE SET switch. "AddXXX" will appear on the display (where "XXX" is one of the four pre-programmed sample sizes). Select the desired sample size by repeatedly pressing the SAMPLE SET switch, or key in an alternative sample size.

IMPORTANT:	The displayed sample size must equal the number of sample items  removed.

6.	Place the sample pieces that were removed back into the container. The scale will momentarily display the percent error/accuracy (If selected in calibration.) and enter the count mode.

- The display will show the total count for the items in the container.

Top-End Counting (Two-Switch method; Sample Removed)

1.	Press ZERO to zero the scale.

2.	Press KEYPAD TARE, key in the tare weight of the container, then press ENT.

3.	Place the container with contents on the platform and press the SAMPLE SET switch. "AddXXX" will appear on the display (where "XXX" is one of the four pre-programmed sample sizes). Select the desired sample size by repeatedly pressing the SAMPLE SET switch, or key in an alternative sample size.

- Remove the displayed sample quantity from the container. The display will show the percent error/accuracy (If selected in calibration.). Press ENT.

5.	The scale will enter the count mode and display the number of parts in the container.

- Return the sample pieces to the container.

Counting by First Determining Error of Count using Automatic Sample Update (Normal Operation; One-Switch Method)

1.	If a container will be used to hold items being counted, place it on the platform.

2.	Press the SAMPLE SET switch. "AddXXX" appears on the display (where "XXX" is one of the pre-programmed sample sizes). Select the desired sample size by repeatedly pressing the SAMPLE SET switch, or key in an alternative sample size.

- Place the sample pieces on the platform simultaneously. The Percent error/accuracy will be momentarily displayed and the scale will enter the counting mode.

4.	To either decrease the percent of error or increase the percent of accuracy, place more items on the platform in the range of 10% to 100% of the original sample quantity. A new piece weight is calculated based on the larger sample size, the new percent of error/accuracy is momentarily displayed and the scale returns to the count mode. This may continue until the desire percent error/accuracy  is achieved.

- Removing the original sample or adding a quantity of parts greater than the original sample will disable the automatic update feature.

Counting by First Determining Error of Count (Two-Switch Method)

1.	If a container will be used to hold the items being counted, place it on the platform.

- Press the SAMPLE SET switch for two seconds. "AddXXX" will appear on the display (where "XXX" is one of the pre-programmed sample sizes). Select the desired sample size by repeatedly pressing the SAMPLE SET button or key in a new sample size.

3.	Place the sample onto the platform all at one time (this should be done if the parts are extremely small) or count them onto the platform one at a time.

- The display will show the percent error/accuracy. Continue to count parts onto the platform, as the sample increases the percent error will decrease and accuracy will increase. When the percent error/accuracy is acceptable key in the quantity of parts on the platform and press ENT.

Counting by First Determining Error of Count with Automatic Sample-To-Bulk Feature (Two-Switch-Method)

- Press the BASE/1 switch to select the bulk base. (BASE 2 light should be lit)

2.	Press the ZERO to zero the bulk base.

- If a container will be used to hold items on the bulk base, place the container on the platform and press ZERO or TARE.

4.	If a container will be used to hold the sample parts, place it onto the sample base at this time.

- Press the SAMPLE SET switch for two seconds. "AddXXX" will appear on the display (where "XXX" is one of the pre-programmed sample sizes). Select the desired sample size by repeatedly pressing the SAMPLE SET switch or key in a new sample size.

6.	Place the sample onto the platform all at once (this should be done if the parts are extremely small) or count them onto the platform one at a time.

- The display will show the percent error/accuracy. If the percent error/accuracy is not acceptable, continue to count parts onto the platform, as the sample increases the percent error will decrease and the percent accuracy will increase. When the percent error/accuracy is acceptable, key in the quantity of parts on the platform and press ENT.

8.	The scale will switch to the bulk base. Items may now be counted on the bulk base.

Negative Counting, Sample Removed (Two Switch Method)

- With nothing on the scale platform, press the TARE to set tare to zero.

2.	Place the container with contents on the platform.

- Press the SAMPLE SET switch. “AddXXX” will appear on the display (where “XXX” is one of the four pre-programmed sample sizes). Select the desired sample size by repeatedly pressing the SAMPLE SET switch, or key in an alternative sample size.

4.	Remove the displayed sample quantity from the container. The display will show the percent error/accuracy (If selected in calibration). Press ENT.

- The scale will enter the count mode and display the number of parts removed from the container as a negative number. Removing pieces increases the negative count.

BI-DIRECTIONAL SERIAL ASCII INTERFACE

INTRODUCTION

The 7600 Scale Serial Ascii Interface is an RS-232C, TTL-compatible, asynchronous serial interface capable of bidirectional transmission of scale data.

This interface features; programmable baud rates, word lengths, stop bits, parity, address number, and a formatted print.

COMMUNICATION CAPABILITIES

The interface provides three basic types of communication with the Model 7600:

1.	Transmission of a predetermined (at set-up time) array of information for formatted printing.  This may be initiated by pressing the PRINT push-button or by sending "SRP"<CR> ("Send Requested Print") command to the unit. The scale is shipped from the factory set up to send the Gross weight when the PRINT button is pressed.

- Transmission and alteration of specific data upon request from an external device.

3.	Alteration of the unit's operating mode upon request from an external device.

- Alteration of scale Calibration, RS232 communications parameters and print formatting information. ( Refer to "Programming and Calibration Manual" for more information. )

RS-232 PIN ASSIGNMENTS AND IMPLEMENTED FUNCTIONS

Connection to the Serial Port is made via a DB-9 female connector found in the access area on the bottom of the scale.

PIN		EIA CODE		 FUNCTION			DIRECTION

2			BB			Transmit Data			  Output

3			BA			Receive Data			  Input

5			AB			Signal Ground		    -

6			CC			Data-Set Ready (DSR)	  Output

7			CB			Clear to Send (CTS)		  Input

8			CA			Request to Send (RTS)	  Output

NOTE:  	All remaining pins are currently unused and unconnected.  For simple terminal usage where the data rate does not exceed either machine's capacity to process it, only pins 2, 3, and 5 need to be connected.

TYPICAL CABLE WIRING:	To connect to an IBM AT or compatible use a cable with one male and one female DB-9 connector with all nine pins connected straight through.

RS-232 COMMUNICATION PARAMETERS

The unit leaves the factory with the following default communication settings :

BAUD		300

WORD LENGTH	7 bits

STOP BITS		1

PARITY		ODD

These settings may be changed using the scale keyboard, or RS232 interface commands.  If the RS232 interface is used, initial communication with the users computer or terminal must be established using the above settings.  To return the RS232 settings to the above default values:

- Unplug the unit power cord or otherwise remove power.

2.	Depress the internal Calibrate button.

- Plug in (or energize) the unit and continue to hold the

calibrate button for 5 seconds.

For information on how to reconfigure the RS-232 parameters, refer to the "Calibration and Programming Manual".

COMMAND FORMATS

The Model 7600 can be controlled from an external device (such as a computer or terminal) by various commands, each three letters long, which represent related English phrases or words.  For example, to tell the scale to zero, type ZRO followed by a carriage return.

When the scale receives command strings, they are first placed in a 125-character buffer.  If many commands are sent to the scale at high baud rates, it is possible to completely fill this buffer and data will be lost.  Be sure to send commands to the scale at a rate which does not exceed 125 characters every 500 ms.

The basic command formats are:

1.  [<add>]<cmd><cr>

2.  <cmd> [<flt>]<cr>

Where <cmd> is a three-letter command, <add> is a scale address number (0-255), <cr> represents a carriage return, and <flt> is a mixed number, the brackets [ ] are used to indicate an optional part of the command.

The following are some examples of command formats:

Command format 1:

SRP<cr>		Send a formatted print

Command format 2:

ITW 13.43<cr>	Instructs scale to set tare weight

to 13.43 in the current unit

Command format 3:

5 SGW<cr>		Instructs the scale with address #5 to

send the Gross weight.

MODEL 7600 SERIES RS-232 SCALE COMMANDS

General Commands

ATW		Acquire Tare Weight

CHK		Initiate self-diagnostics CHecK

LCK		LoCK out keypad

PON		Power ON

POF		Power OFf

RES		RESet, clears tare weight and count information

SCM		Select Count Mode

SSS		Select Sample Size

SWM		Select Weight Mode

UCK		UnloCK keypad

UNP		UNit Primary

UNS		UNit Secondary

ZRO		ZeRO scale

Commands Which Enter Information into the Unit

IBA [ FLOATING POINT NUMBER ]		Input BAse number (with option)

IPW [ FLOATING POINT NUMBER ]		Input Piece Weight

ITW [ FLOATING POINT NUMBER ]		Input Tare Weight

IID [ FLOATING POINT NUMBER ]		Input ID number

Commands Which Request Information

SAI		Send All Information

SAO		Send Abbreviated Output (same as SRP)

SBA		Send BAse in use (with second base option)

SCI		Send Configuration Information

SCO		Send COunt

SDT		Send DaTe

SGW		Send Gross Weight

SID		Send Part ID

SMI		Send Metrological Information

SNW		Send Net Weight

SPC		Send Print Codes

SPR		Send PeRcent error or accuracy

SPW		Send Piece Weight

SRP		Send Requested Print

SSZ		Send Sample siZe

STM		Send TiMe

STW		Send Tare Weight

SVN		Send Software Version Number

Calibrate and Configure Commands

CAL		CALibrate - Same as pressing SW4 button.

CFC		ConFigure Communication - To set baud rate etc.

CFP		ConFigure Print codes - To enter print codes.

CLE		CaLibration End - To save calibration data.

CLM		CaLibrate Modes - To set Auto zero, Count,

PcWt/1000

CLP		CaLibrate Primary - To set-up primary calibra-

tion data.

CLS		CaLibrate Secondary - To set-up secondary cali-

bration data.

CLU		CaLibration Unstable - To set-up print

stable/unstable.

CLW		CaLibration Weight - To calibrate with weights.

NOTES: 	All commands and parameters must be separated by spaces. The entire command string must be terminated with a carriage return.

All calibrate and configure commands are further explained in the "Calibration and Programming Manual".

SAMPLE COMPUTER PROGRAM

The 7600 series RS-232 interface may be tested by connecting it to a terminal with an RS-232 port, or a computer with an RS-232 and a terminal program. If a terminal program is not available the following program is provided for an IBM PC or compatible using BASICA, MBASIC or QuickBASIC.

10 '  Terminal Program written for BASIC Programming Language

20 '   Provided by Pennsylvania Scale Company

30 '

40   DEFINT A-Z

50   FALSE = 0: TRUE = NOT (FALSE)

60   KEY OFF: CLS : LOCATE 25, 1

70   PRINT SPACE$(20); "Press Alt-X to Exit Terminal Program"

80   LOCATE 1, 1

90   OPEN "Com1:300,o,7,1" FOR RANDOM AS #1 'Handshaking Enabled

100 OPEN "test.dta" FOR OUTPUT AS #2

110 WHILE NOT (QUIT)

120  KEYBOARDINPUT$ = INKEY$

130  IF KEYBOARDINPUT$ = CHR$(0) + CHR$(45) THEN QUIT = TRUE

140  IF KEYBOARDINPUT$ <> " " THEN PRINT #1, KEYBOARDINPUT$;

150  IF LOC(1) > 0 THEN

160  SCALEINPUT$ = INPUT$ (LOC(1), #1)

170  PRINT SCALEINPUT$;	'Writes data from scale to screen

180  LPRINT SCALEINPUT$;	'Writes data from scale to printer

190  PRINT #2, SCALEINPUT$;	'Writes data from scale to file

200  END IF

210  WEND

220  END

NOTES:	Omit line 180 if data is not to be sent to the printer.

Omit lines 100 and 190 if data is not to be written to a file.

In line 100, the word "OUTPUT" opens the file "test.dta", deleting any data that was in the file before this program was started. If you want to preserve the contents on the file "test.dta" from one session to another, replace the word OUTPUT with APPEND on line 100.

NOTES: (cont’d.)

Line 90 opens the serial communications port, using 300 baud, Odd parity, 7 data bits, and 1 stop bit. These represent the default settings for the scale, other entries may be used if the scale settings have been changed. Valid entries for baud rate are 300, 600, 2400, 4800, 9600, or 19200. Valid entries for parity are N (none), E (even), or O (odd). Valid entries for data bits are 7 or 8. Valid entries for stop bits are 1 or 2. Com2 may be used if  the scale is attached to serial communications port #2.

To eliminate the need for attaching the handshaking lines in the cable between the scale and the computer, change line 90 to read:

90  OPEN "Com1:300,o,7,1,RS,DS0" FOR RANDOM AS #1

'Handshaking Disabled

ERROR MESSAGE LIST

During Diagnostics:

Err1.CA			Scale Calibration

Err2.00 to Err2.63		EEPROM Write Failure

Err2.rS			Serial Interface Setup

Err3.Hd			Deadload too High

Err3.nS			No Sense Line Connected

Err3.-d			Negative Deadload

Err4.Sh			Load Cell Excitation Shorted

Err6.-1 to Err6.-4		Switch SW1-SW4 Shorted

Err7.nr			Defective NOV RAM

Err8.00 to Err8.FF		RAM Test Error

Err9.CS			Eprom Checksum Error

All Other Times:

Message on Display          	Explanation			Message From ASCII Port

Err2.00 to Err2.FF		EEPROM read error		"Err2.? EEPROM"

Err 10			Count too large		none

Err 11			Keyboard Entry Error	none

Err 13			Negative number less		none

than - 99,999 (Will not

fit on display.)

Err 30			Illegal Zero			none

Err 31			Illegal Tare			none

Err 40			Load Cell error		none

Err 80			RS - 232 Input Error		none

Err 81			RS - 232 Input Error		none

Err 82			RS - 232 Overflow		none

Err Cnf			Config load error on startup	"Err 85 Reset to 300 baud"

Err xx*			Undefined (Consult Factory)	none

ULULUL			Load Cell Underload		"Err 41"

OLOLOL			Load Cell Overload		"Err 42"

CalErr			RS - 232 Cal Error		"Calibration Command Error"

none			Bad Print Code		"Err 83 Print Code"

none			No End Code			"Err 84 No Code 99"

- Where "xx"  equals any number or character not shown on the above list.